Training Notes

Primarily biological information and testing results

18/11/2025:

- Assembled the main Pioreactor vessel

- Prepared the electrolytically competent media for Cupriavidus metallidurans cultivation based on the Sydow optimised media by Martin (removal of all phosphates)

  - 25ml of meso-nutrients; 1.25ul of trace elements, 85ul of iron sulphate

    - For heterotrophic variant add 10mM (final concentration) of gluconic acid

  - https://github.com/amy-bo/electroPioreactor/tree/main/Media

  - The media was filter sterilised using a 0.2 micron filter due to sedimentation and colour change of trace elements after autoclaving

    - At that moment, we did not have Manganese sulphate Monohydrate, instead we used 0.317 g of Manganese II Sulphate Tetrahydrate (molar  quantity ends up the same)

  - Inoculated 2x 25ml of media at 30C for 17 hours, heterotrophically

    - Original liquid culture OD of 0.614

    - 1st culture final OD of 0.001

    - 2nd culture final OD of 0.02

19/11/2025:

- Finish the pioreactor set-up

- Overnight ODs

  - Starting point 0.001 = 0.160 OD

  - Starting point 0.02 = 1.015 OD

- Electrical analysis of the system

|  | Pioreactor 1 | Pioreactor 2 | Pioreactor 3 | Pioreactor 4 |
| --- | --- | --- | --- | --- |
| Voltage | 3.38 | 3.41 | 3.36 | 3.28 |
| current | 0.012 | 0.012 | 0.012 | 0.012 |

- Inoculated the pioreactor with 15ml of autotrophic media (without gluconic acid) to final ODs of 0.001 (pioreactor 1 and 2) and 0.02(pioreactor 3 and 4)

20/11/2025:

- No growth due to faulty OD sensors. Possible causes are:

  - Suboptimal media composition

  - High voltage

  - Inconsistent temperature maintenance

  - OD sensor being blocked by an electrode

- Further pioreactor training continued, such as setting up media and waste tubes and how to calibrate it.

21/11/2025:

- Carried out a conductivity test of the autotrophic media

  - 6.48MS at 18.8C

  - 6.55 MS at 19.1C

| Percent output (as shown on the software) | Voltage |
| --- | --- |
| 1% | 2.510 |
| 2% | 2.619 |
| 3% | 2.720 |
| 4% | 2.8 |
| 5% | 2.913 |
| 6% | 2.961 |
| 7% | 2.97 |

- Noticed the discoloration on the cathode, near the CO2 sparging tube