## Model
electroPioreactor - CO₂ dosing & O₂ management model
AMYBO aseptic electro-bioreactor - hydrogen-oxidising bacteria (C. necator) grown on electrolytic H₂/O₂ + dosed CO₂. UK English, SI. PHASE 1 (refined). Rule: every cell is a number, text, or a formula; formulae carry no embedded data values - each datum lives in its own named, sourced cell (only self-evident unit/stoichiometric integers appear inline). Invalid selector entries return an error by design. Sensitivity analysis deferred.
Confidence legend (source-cell colour):
Verified / handbook / defined / standard
Literature-supported (~90%)
Design assumption (~70%)
Weak / coarse estimate (~50%)
Guess - replace when known
DATA GAP - measure, do not invent
0. SELECTORS (each picker has exactly two valid values; anything else errors)
Parameter | Name | Unit | Value | Source / assumption
Reactor version | rx_ver | - | AEP0.1.1 | Active build. Choose from the two valid values below. Dropdown.
Sparger type | sparger | - | Sintered | CO₂ delivery route. Choose from the two valid values below. Dropdown.
Sinter porosity grade (if Sintered) | por_grade | 0-5 | 0 | ISO 4793 / DURAN class; 0 = coarsest. Dropdown.
valid version 1 | ver_name_1 | - | AEP0.1.1 | Current build (20 mL Pioreactor vial, assumed).
valid version 2 | ver_name_2 | - | AEP0.2 | Planned build (40 mL Pioreactor vial).
valid sparger 1 | spg_name_1 | - | Tube | Bare 1/8" OD x 1/16" ID Flexelene 135C tube orifice.
valid sparger 2 | spg_name_2 | - | Sintered | 10 mm sintered-glass disc annulus on a tubular anode.
1A. REACTOR VERSION DATA (raw per-build values - each its own sourced cell)
Parameter | Name | Unit | Value | Source / assumption
Vial outer diameter - AEP0.1.1 | vial_OD_1 | mm | 27.48 | Pioreactor 20 mL vial, measured OD.
Vial outer diameter - AEP0.2 | vial_OD_2 | mm | 28 | 40 mL vial OD - placeholder, measure.
Internal usable depth - AEP0.1.1 | D_int_1 | mm | 55 | 20 mL vial total height 57.4 mm; usable depth est. Measure.
Internal usable depth - AEP0.2 | D_int_2 | mm | 95 | 40 mL vial - placeholder, measure.
Recommended max working volume - AEP0.1.1 | Vmax_1 | mL | 16 | Pioreactor docs: 8-16 mL for 20 mL vial; using max as level datum.
Recommended max working volume - AEP0.2 | Vmax_2 | mL | 25 | Pioreactor: 40 mL vial max 25 mL.
Total internal vial volume - AEP0.1.1 | Vtot_1 | mL | 24 | To cap underside (> nominal). Placeholder - measure (water to fill).
Total internal vial volume - AEP0.2 | Vtot_2 | mL | 42 | Placeholder - measure.
1B. ACTIVE REACTOR GEOMETRY (selected build; level-matched to bare max volume)
Parameter | Name | Unit | Value | Source / assumption
Vial outer diameter | vial_OD | mm | 27.48 | Selected by rx_ver; invalid version returns an error.
Vial wall thickness | vial_wall | mm | 1.1 | Borosilicate vial; Pioreactor publish neither wall nor internal diameter, so wall estimated ~1.1 mm (typical for this glass-vial class).
Vial internal diameter | vial_ID | mm | 25.28 | OD minus two wall thicknesses.
Internal cross-sectional area | A_x | mm² | 501.931 | Uniform-cylinder body approximation (neck excluded - see headspace).
Internal usable depth | D_int | mm | 55 | Selected by rx_ver.
Recommended max working volume (datum) | V_max | mL | 16 | Selected by rx_ver.
Total internal vial volume | V_vial_total | mL | 24 | Selected by rx_ver.
Bare liquid level (datum height) | h_datum | mm | 31.8769 | Height the bare max volume reaches with no inserts.
Electrode rod diameter | rod_d | mm | 6 | 6 mm rod electrodes.
Number of electrodes | rod_n | - | 2 | Anode + cathode.
Electrode tip clearance above base | elec_clear | mm | 22 | Held CONSTANT across builds (the OD-measurement clear zone). 22 mm = the 33 mm insertion used in AEP0.1.1 (D_int 55). Raise it (e.g. to 25) if OD is disrupted during electrolysis. Input.
Electrode insertion depth (derived) | elec_ins | mm | 33 | Cap underside -> tip. Increases for the taller AEP0.2 to keep clearance constant.
Sparge tube OD | spg_OD | mm | 3.175 | 1/8" OD Flexelene 135C.
Sparge tube ID | spg_ID | mm | 1.5875 | 1/16" ID.
Sparge release height (= anode tip) | spg_tip_h | mm | 22 | CO₂ released near anode tip.
Efflux tube OD | eff_OD | mm | 3.175 | 1/8" OD Flexelene 135C.
Efflux tube ID | eff_ID | mm | 1.5875 | 1/16" ID.
Electrode submerged length (@datum) | elec_sub_L | mm | 9.8769 | For displacement bookkeeping.
Electrode displaced volume | elec_disp | mL | 0.55853 | Submerged rod volume.
Sparge tube displaced volume | spg_disp | mL | 0.078198 | Submerged sparge-tube volume.
Efflux tube submerged length | eff_sub_L | mm | 2.3813 | 45 deg cut; level held at top of the ID opening, so wall is wetted from tip up to (OD+ID)/2.
Efflux tube displaced volume | eff_disp | mL | 0.01414 | Annular wall only (bore is open to withdrawal). 45 deg approximation.
Total displaced volume | disp_tot | mL | 0.65086 | Sum of submerged inserts.
LIQUID VOLUME TO CHARGE | V_charge | mL | 15 | Integer mL to charge so the level matches the bare datum despite inserts.
Resulting liquid level | h_actual | mm | 31.1813 | Level achieved by the integer charge (displacement taken at datum; sub-mm error).
Sparge depth / bubble path length | sparge_depth | mm | 9.1813 | Release point to surface.
Gas-liquid interfacial area | interface_A | mm² | 429.5477 | Free surface = cross-section minus inserts piercing it.
Extra headspace tube OD | xtube_OD | mm | 3.175 | 3 further 1/8" tubes (gas-out, etc.) through the cap.
Number of extra headspace tubes | xtube_n | - | 3 | Protrude through the cap into headspace.
Extra tube protrusion | xtube_pro | mm | 5 | Each protrudes ~5 mm below cap (45 deg cut). Input.
Total in-vial insert volume | V_inserts | mL | 2.4347 | All inserts inside the vial (submerged + headspace): electrodes & sparge full length, efflux to surface, 3 extra tubes.
Headspace gas volume | headspace_V | mL | 6.5653 | Total vial volume - liquid charged - inserts (captures neck narrowing).
2. ELECTROLYSIS (Gerrit's Law: constant current from LED intensity)
Parameter | Name | Unit | Value | Source / assumption
LED intensity setpoint (Pioreactor LED_D) | intensity | % | 3 | Electrode drive via LED_D channel. Input.
Gerrit's Law - slope | gerrit_slope | mA/% | 1.03 | Empirical linear fit of current vs LED intensity (constant-current: independent of electrolyte, volume, electrode coverage).
Gerrit's Law - intercept | gerrit_int | mA | 2.6 | Intercept of the same fit.
Gerrit's Law - min valid intensity | gerrit_min | % | 3 | Fit validated from 3%.
Gerrit's Law - max valid intensity | gerrit_max | % | 25 | Fit validated to 25%; above this constant-current may fail as cell V nears the 5 V rail.
Faraday constant | F_const | C/mol | 96485.3321 | NIST CODATA (exact, SI 2019).
Electrons per H₂ | z_e_H2 | - | 2 | 2H⁺ + 2e⁻ -> H₂.
Electrons per O₂ | z_e_O2 | - | 4 | 2H₂O -> O₂ + 4H⁺ + 4e⁻.
Faradaic efficiency | etaF | - | 1 | Clean water electrolysis FE approaches 100% (Nat. Commun. 2022). But in this neutral-pH cell, cathodic O₂ reduction competes with H₂ evolution, so true H₂ FE is likely <1 and UNMEASURED. 1.0 = optimistic upper bound; measure by gas collection.
Electrolysis current (Gerrit's Law) | I_app | A | 0.00569 | I(mA) = slope x intensity + intercept; /1000 -> A.
Current-law validity flag | I_valid | - | OK | Linear fit only validated over the stated range.
H₂ generation rate | rH2_gen | mol/h | 0.00010615 | Faraday: I x etaF /(z F).
O₂ generation rate | rO2_gen | mol/h | 5.3075e-05 | H₂:O₂ = 2:1.
3. BIOLOGY (cells consume 100% of evolved H₂; O₂ & CO₂ pro-rata)
Parameter | Name | Unit | Value | Source / assumption
Gas requirement ratio - H₂ | bio_H2 | - | 6 | H₂:O₂:CO₂ consumption ratio (starting 6:2:1). NB literature recommends a 7:2:1 FEED mix - feed is not consumption.
Gas requirement ratio - O₂ | bio_O2 | - | 2 | As above.
Gas requirement ratio - CO₂ | bio_CO2 | - | 1 | As above.
H₂ consumed | H2_cons | mol/h | 0.00010615 | Assumption: 100% of evolved H₂ consumed.
O₂ consumed | O2_cons | mol/h | 3.5384e-05 | H₂ consumed x (O₂/H₂ ratio).
CO₂ consumed (carbon fixation) | CO2_cons | mol/h | 1.7692e-05 | H₂ consumed x (CO₂/H₂ ratio).
EXCESS O₂ to remove | O2_excess | mol/h | 1.7692e-05 | Electrolysis makes more O₂ than the cells use; surplus must be stripped.
4. O₂ CONSTRAINT (growth ceiling; cathodic competition is the real driver)
Parameter | Name | Unit | Value | Source / assumption
Temperature | T_C | °C | 30 | C. necator optimum ~30 °C. Input.
Temperature (absolute) | T_K | K | 303.15 | °C + 273.15.
Surface pressure | P_atm | Pa | 101325 | Atmospheric at the liquid surface (assumed).
Standard atmosphere (atm->Pa) | Pa_per_atm | Pa/atm | 101325 | Defined: 1 atm = 101 325 Pa exactly.
Henry solubility O₂ (ref, 298.15 K) | H_O2ref | mol/m³/Pa | 1.3e-05 | Sander (2023) compilation (supersedes 2015), value from Burkholder et al. 2019.
Henry T-dependence O₂ | H_O2T | K | 1500 | Sander (2023): d ln(Hcp)/d(1/T).
Reference temperature | T_ref | K | 298.15 | Standard reference for Henry constants.
Henry solubility O₂ at T | H_O2_T | mol/m³/Pa | 1.1965e-05 | van 't Hoff correction to reactor T.
O₂ growth-inhibition ceiling (gas) | O2_ceil_atm | atm | 0.3 | Whole-cell growth inhibited above ~0.30 atm O₂ - NOT a hydrogenase limit (the [NiFe] enzyme is O₂-tolerant).
hydrogenase note |  | - | (O₂-tolerant) | C. necator membrane-bound [NiFe]-hydrogenase tolerates O₂, so the enzyme is not the binding constraint.
O₂ ceiling (partial pressure) | O2_ceil_Pa | Pa | 30397.5 | Ceiling in Pa.
Dissolved-O₂ ceiling | O2_ceil_C | mol/m³ | 0.3637 | Henry's law at the gas ceiling.
Dissolved-O₂ ceiling | O2_ceil_uM | µM | 363.7005 | mol/m³ = mmol/L; x1000 = µM. Practical aim is to MINIMISE dissolved O₂ (cathodic O₂ reduction steals H₂-evolution current); floor set by stripping capacity.
5. CO₂ DOSING (SodaStream + regulator + solenoid + needle valve)
Parameter | Name | Unit | Value | Source / assumption
CO₂ volumetric flow rate | Q_CO2 | mL/min | 10 | Needle-valve setting at ~atmospheric surface. Input.
Sparge duration per pulse | spg_dur | s | 1 | Solenoid open time. Input.
Sparge interval | spg_int | min | 1 | Time between pulse starts. Input.
Min reliable pulse (solenoid floor) | pulse_floor | s | 0.5 | Smallest dependable actuation. Provisional - characterise.
Ideal gas constant | R_gas | J/mol/K | 8.3145 | NIST CODATA (exact, SI 2019).
Molar mass CO₂ | M_CO2 | g/mol | 44.0095 | NIST WebBook.
CO₂ molar flow during sparge | nCO2_rate | mol/s | 6.7e-06 | Ideal gas n=PV/RT at surface conditions.
CO₂ per pulse | CO2_pulse | mol | 6.7e-06 | Molar flow x duration.
Pulses per hour | pulses_h | 1/h | 60 | 60 / interval(min).
Pulses per day | pulses_d | 1/d | 1440
Average CO₂ supply rate | CO2_supply | mol/h | 0.000402 | Time-averaged.
CO₂ supply : carbon demand | CO2_sd_ratio | × | 22.7223 | >1 = carbon non-limiting (before stripping use).
Sparge duty cycle | duty | - | 0.016667 | Fraction of time sparging.
6. CO₂ BUBBLE SIZE (orifice vs sinter)
Parameter | Name | Unit | Value | Source / assumption
Surface tension (water, T) | sigma | N/m | 0.0712 | IAPWS R1-76(2014) formulation, evaluated at 30 °C.
Liquid density (water, T) | rho_L | kg/m³ | 995.65 | IAPWS-95, evaluated at 30 °C.
Gravitational acceleration | g_const | m/s² | 9.8066 | Standard gravity (defined).
O₂ diffusivity in water (T) | D_O2 | m²/s | 2.4e-09 | Literature ~2.4e-9 at 30 °C (Han & Bartels 1996). Confirm.
Pore size - grade 0 | por0_um | µm | 205 | DURAN P0 160-250 µm, midpoint. ISO 4793 (paywalled) underlies; DURAN table is free.
Pore size - grade 1 | por1_um | µm | 130 | DURAN P1 100-160 µm, midpoint.
Pore size - grade 2 | por2_um | µm | 70 | DURAN P2 40-100 µm, midpoint.
Pore size - grade 3 | por3_um | µm | 28 | DURAN P3 16-40 µm, midpoint.
Pore size - grade 4 | por4_um | µm | 13 | DURAN P4 10-16 µm, midpoint.
Pore size - grade 5 | por5_um | µm | 1.3 | DURAN P5 1.0-1.6 µm, midpoint.
Selected sinter pore diameter | por_pore | m | 0.000205 | Picked by por_grade.
Tube orifice diameter | tube_d_m | m | 0.0015875 | Tube ID in metres.
Effective orifice/pore diameter | d_orifice | m | 0.000205 | Tube ID or sinter pore; invalid sparger returns an error.
Bubble detachment diameter | d_bubble | m | 0.0020777 | Static low-flow force balance (buoyancy vs surface tension). Flow-dependence not modelled.
Bubble diameter (display) | d_bubble_mm | mm | 2.0777
Mendelson coefficient a | mend_a | - | 2.14 | Mendelson (1967) wave-theory bubble rise correlation (AIChE J, paywalled).
Mendelson coefficient b | mend_b | - | 0.505 | Mendelson (1967), as above.
Bubble rise velocity | u_rise | m/s | 0.28973 | Mendelson correlation, mm-scale bubbles.
7. O₂ STRIPPING CAPACITY (kLa estimate - low confidence)
Parameter | Name | Unit | Value | Source / assumption
Superficial gas velocity (sparge on) | u_sg | m/s | 0.00033205 | Gas flow / liquid cross-section.
Gas holdup (sparge on) | holdup | - | 0.0011461 | ε ≈ superficial gas velocity / rise velocity.
Interfacial area for transfer | a_int | 1/m | 3.3096 | a = 6ε/d_b (spheres).
Bubble-liquid contact time | t_contact | s | 0.0071712 | d_b / rise velocity.
Liquid-film coefficient kL | kL | m/s | 0.00065278 | Higbie penetration theory.
kLa during sparge | kLa_sparge | 1/s | 0.0021604 | kL x a.
kLa time-averaged | kLa_avg | 1/s | 3.6007e-05 | x duty cycle.
O₂ strip rate at ceiling (sparge on) | strip_sparge | mol/h | 4.243e-05 | Max while sparging, bulk O₂ at ceiling, CO₂ bubbles ~O₂-free.
O₂ strip rate time-averaged | strip_avg | mol/h | 7.0717e-07 | x duty cycle.
Strip capacity : excess-O₂ | strip_ratio | × | 0.039972 | >=1 = dissolved-gas stripping holds the ceiling; <1 = relies on bubble knock-off or DO rises.
Required time-avg kLa | kLa_req | 1/s | 0.00090081 | kLa that would exactly strip the excess at the ceiling.
8. MEDIA CARRYOVER CEILING (vent-filter clog risk)
Parameter | Name | Unit | Value | Source / assumption
Entrainment velocity ceiling (rule of thumb) | u_g_max | m/s | 0.05 | Provisional superficial-velocity ceiling above which droplet carryover to the hydrophilic vent filters becomes likely. Replace with measured value.
Carryover flag | carry_flag | - | OK | Operating gas velocity vs ceiling.
9. DOSING-SCHEDULE FEASIBILITY
Parameter | Name | Unit | Value | Source / assumption
O₂ accumulated per interval | O2_accum | mol | 2.9486e-07 | Excess O₂ generated between pulses.
O₂ removable per pulse | O2_strip_pulse | mol | 1.1786e-08 | O₂ a pulse strips (at ceiling).
Per-pulse balance | sched_bal | × | 0.039972 | >=1 = each pulse clears the interval's O₂ by stripping alone.
Duration floor check | dur_ok | - | OK | Pulse must exceed reliable actuation time.
Schedule verdict | sched_verdict | - | Stripping alone insufficient - rely on bubble knock-off / raise flow or duty / accept higher DO | Synthesised recommendation (Phase 1 logic).